[image: image66.png]Ministry for the

Environment

Manatu Mo Te Taiao

&,






GEMS/AMIS Air Quality Monitoring Programme Annual Report 2010
Prepared for the Ministry for the Environment
by

	[image: image68.jpg]e NO2 pg/m? === NO pg/m*




	Laboratory Services

Air Quality Department


Disclaimer

This report or document (the report) is given by Watercare Services Limited (Watercare) solely for the benefit of the Ministry for the Environment (MfE) as defined in the contract or terms and conditions between Watercare and MfE, and is strictly subject to the conditions laid out in that contract or terms and conditions. This report may not be reproduced, except in full.
Neither Watercare nor any of its employees makes any warranty, express or implied, or assumes any legal liability or responsibility for use of the report or its contents by any other person or organisation.

Watercare Services Ltd
52 Aintree Avenue

Airport Oaks

PO Box 107 028

Airport Oaks

Auckland

Ph 09 539 7600

Fax 09 539 7601

Prepared by:
Reviewed by:

Kathleen McLeod
Jonathan Taylor

Senior Air Quality Technician
Air Quality Technician
and Key Technical Person
and Key Technical Person

This report may be cited as:

Ministry for the Environment. 2011. GEMS/AMIS Air Quality Monitoring Programme Annual Report 2010. Wellington: Ministry for the Environment.

Published in December 2011 by the
Ministry for the Environment
Manatū Mō Te Taiao
PO Box 10362, Wellington 6143, New Zealand

ISBN: 978-0-478-37274-8
Publication number: ME 1085
© Crown copyright New Zealand 2011
This document is available on the Ministry for the Environment’s website:
www.mfe.govt.nz
[image: image2.png]Ministry for the

Environment

Manatu Mo Te Taiao

&,





Contents

11
Summary

2
Introduction
3
3
Air Pollutants Monitored
4
3.1
Carbon monoxide
4
3.2
Nitrogen oxides
4
3.3
Sulfur dioxide
5
3.4
Volatile organic compounds
5
3.5
Particulate matter
5
3.5.1
Particulate matter as PM10
6
3.5.2
Total suspended particulate
6
3.6
Lead
6
4
Ambient Air Quality Guidelines and Standards
7
5
Monitoring Sites
8
5.1
Site descriptions
8
5.1.1
MfE Gavin Street, Penrose, Auckland – Site AKL009
8
5.1.2
MfE Greers Road, Burnside, Christchurch – Site CAN002
9
5.1.3
MfE Coles Place, St Albans, Christchurch – Site CAN003
10
5.2
Air pollutants monitored under the GEMS/AMIS programme
11
6
Methods
12
6.1
Quality assurance
12
6.2
Analytical methods
13
6.2.1
Carbon monoxide
13
6.2.2
Nitrogen oxides
13
6.2.3
Sulfur dioxide
13
6.2.4
Volatile organic compounds (VOCs)
13
6.2.5
Particulate matter as PM10
14
6.2.6
Total suspended particulate matter (TSP)
14
6.2.7
Lead
14
7
Results and Discussion
15
7.1
Site performance and quality assurance
15
7.2
Carbon monoxide (CO) 2010
16
7.3
Nitrogen oxides (NO2 and NO) 2010
16
7.4
Sulfur dioxide (SO2) 2010
16
7.5
Volatile organic compounds (VOC) January–December 2010
17
7.6
Particulate matter (PM10) 2010
22
7.7
Total suspended particulates (TSP) 2010
23
7.8
Lead (Pb) June–August 2010
26
7.9
Analysis of exceedences
41
7.9.1
Exceedences at Greers Road, Burnside
41


Figures

22Figure 1:
MfE Benzene annual average 2010

Figure 2:
Auckland TSP 7-day average January–December 2010
24
Figure 3:
Auckland TSP 7-day average 1994–2010
24
Figure 4:
Christchurch TSP 7-day average January–December 2010
25
Figure 5:
Christchurch TSP 7-day average 2002–2010
25
Figure 6:
MfE lead 3-month average results 1996–2010
26
Figure 7:
MfE Burnside CO 1-hour fixed average January–December 2010
27
Figure 8:
MfE Burnside CO 1-hour fixed average 2003–2010
27
Figure 9:
MfE Burnside CO 8-hour rolling average January–December 2010
27
Figure 10:
MfE Burnside CO 8-hour rolling average 2003–2010
28
Figure 11:
MfE Burnside CO annual average 2003–2010
28
Figure 12:
MfE Penrose NO2 1-hour fixed average January–December 2010
28
Figure 13:
MfE Penrose NO2 1-hour fixed average 1997–2010
29
Figure 14:
MfE Penrose NO2 24-hour fixed average January–December 2010
29
Figure 15:
MfE Penrose NO2 24-hour fixed average 1 January 1997–2010
29
Figure 16:
MfE Penrose NO2 annual average 1997–2010
30
Figure 17:
MfE Penrose NO2 and NO 1-hour fixed average January–December 2010
30
Figure 18:
MfE Penrose NO2 and NO 24-hour fixed average January–December 2010
30
Figure 19:
MfE Penrose NO 1-hour fixed average 1997–2010
31
Figure 20:
MfE Penrose NO 24-hour fixed average 1997–2010
31
Figure 21:
MfE Penrose NO annual average 1997–2010
31
Figure 22:
MfE Burnside NO2 1-hour fixed average January–December 2010
32
Figure 23:
MfE Burnside NO2 1-hour fixed average 2003–2010
32
Figure 24:
MfE Burnside NO2 24-hour fixed average January–December 2010
32
Figure 25:
MfE Burnside NO2 24-hour fixed average 2003–2010
33
Figure 26:
MfE Burnside NO2 annual average 2003–2010
33
Figure 27:
MfE Burnside NO2 and NO 1-hour fixed average January–December 2010
33
Figure 28:
MfE Burnside NO2 and NO 24-hour fixed average January–December 2010
34
Figure 29:
MfE Burnside NO 1-hour fixed average 2003–2010
34
Figure 30:
MfE Burnside NO 24-hour fixed average 2003–2010
34
Figure 31:
MfE Burnside NO annual average 2003–2010
35
Figure 32:
MfE Penrose SO2 1-hour fixed average January–December 2010
35
Figure 33:
MfE Penrose SO2 1-hour fixed average 2003–2010
35
Figure 34:
MfE Penrose SO2 24-hour fixed average January–December 2010
36
Figure 35:
MfE Penrose SO2 24-hour fixed average 2003–2010
36
Figure 36:
MfE Penrose SO2 annual average 2003–2010
36
Figure 37:
MfE Burnside SO2 1-hour fixed average January–December 2010
37
Figure 38:
MfE Burnside SO2 1-hour fixed average 2003–2010
37
Figure 39:
MfE Burnside SO2 24-hour fixed average January–December 2010
37
Figure 40:
MfE Burnside SO2 24-hour fixed average 2003–2010
38
Figure 41:
MfE Burnside SO2 annual average 2003–2010
38
Figure 42:
MfE Penrose PM10 24-hour fixed average January–December 2010
38
Figure 43:
MfE Penrose PM10 24-hour fixed average 2003–2010
39
Figure 44:
MfE Penrose PM10 annual average 2003–2010
39
Figure 45:
MfE Burnside PM10 24-hour fixed average January–December 2010
39
Figure 46:
MfE Burnside PM10 24-hour fixed average 2003–2010
40
Figure 47:
MfE Burnside PM10 annual average 2003–2010
40
Figure 48:
MfE Burnside, PM10 vs time of day
43
Figure 49:
MfE Burnside, PM10 vs 1.5 m temperature
43
Figure 50:
MfE Burnside, PM10 vs 10 m temperature
43
Figure 51:
MfE Burnside, PM10 vs temperature diff 10–1.5 m
44
Figure 52:
MfE Burnside, PM10 vs. relative humidity
44
Figure 53:
MfE Burnside, PM10 vs wind direction
44
Figure 54:
MfE Burnside, PM10 vs wind speed
45
Figure 55:
MfE Burnside, PM10 vs NO
45
Figure 56:
MfE Burnside, PM10 vs CO
45
Figure 57:
MfE Burnside, PM10 vs SO2
46


Tables

7Table 1:
National environmental standards and guidelines

Table 2:
Air pollutants monitored under the GEMS/AMIS programme 2010
11
Table 3:
Percentage valid and capture data 2010
15
Table 4:
VOC results (January–March 2010)
18
Table 5:
VOC results (April–June 2010)
19
Table 6:
VOC results (July–September 2010)
20
Table 7:
VOC results (October–December 2010)
21
Table 8:
Greers Road, Burnside PM10 exceedences of the daily NES 2010
23


1
Summary
This report contains the 2010 annual ambient air quality data set, measured from Auckland and Christchurch. The parameters form New Zealand’s contribution to the Global Environmental Monitoring System (GEMS) run by the World Health Organization (WHO). The WHO stopped aggregating international air quality information from GEMS in 2006.
The GEMS monitoring sites were established to measure key air pollutants associated with adverse effects on people’s health and wellbeing. The GEMS sites have provided monitoring data for Auckland and Christchurch, and have recorded trends in pollution levels in two of New Zealand’s largest cities.

The Gavin Street, Penrose site in Auckland is representative of primarily vehicle and industrial emissions. The Greers Road, Burnside site in Christchurch is located within a residential area and represents emissions primarily from domestic properties. Air quality monitoring conducted by MfE under the GEMS programme is scheduled to end in June 2011 for the Auckland site and December 2010 for the Christchurch sites.
Air pollutants arise from a number of different sources. The pollutants monitored at the GEMS sites include:

· particulate matter (PM10) – principally from domestic heating, vehicle emissions and industrial processes (as well as natural processes)
· nitrogen oxides (NO and NO2) – mostly from motor vehicle emissions

· sulfur dioxide (SO2) – from the burning of fossil fuels, particularly coal and oil
· carbon monoxide (CO) – from incomplete combustion of carbon containing fuels, especially from motor vehicles
· total suspended particulates (TSP) – from various chemical and physical processes including all particulates less than 100 μm in aerodynamic diameter
· volatile organic compounds (VOCs) – organic chemicals, such as hydrocarbons, that are closely tied to vehicle emissions and many industrial processes
· lead (Pb) – historically from petrol but since the removal of lead in petrol from 1996, levels have declined in New Zealand.
Ambient air quality standards for CO, NO2, SO2, PM10 and ozone were promulgated in October 2004. The purpose of these standards is to provide a guaranteed level of protection for the health of all New Zealanders.
 These standards are based largely on the ambient air quality guidelines (AAQG) developed in 1994 and revised in 2002. Guideline levels for pollutants (and averaging periods) not covered by the standards still apply.
All air pollutants monitored at the MfE sites (with the exception of Greers Road, Burnside) were below the National Environmental Standards (NES) for Air Quality as well as the ambient air quality guidelines (AAQG) for all averaging periods. The Greers Road, Burnside site had eight exceedences of the 24-hour standard for PM10, mainly during the colder months from June to July 2010. These exceedences were most likely due to home heating emissions. The NES for air quality allows for one exceedence of the PM10 24-hour threshold per year before the standard is breached, meaning the Greers Road, Burnside site breached the PM10 standard on seven occasions.
2
Introduction

This report presents the 2010 ambient air quality data set from the GEMS monitoring programme for Auckland and Christchurch, New Zealand. The monitoring is conducted by Watercare Services Limited (Watercare), on behalf of the Ministry for the Environment (MfE).

This data has historically formed New Zealand’s contribution to the World Health Organization’s (WHO) Global Environmental Monitoring System / Air Pollution Programme (GEMS/AIR) which began in 1973.

In 1996, WHO developed the Air Management Information System (AMIS) the successor to GEMS/AIR. The objective of AMIS is to transfer information on air pollutant concentrations and air quality management between countries. The WHO ceased aggregating international air quality information in 2006. The monitoring from the AMIS programme is used by MfE to support air quality monitoring and management in Auckland and Christchurch.
The Auckland site is located in the industrial area of Penrose to the south-east of the city centre. Air quality monitoring has been conducted in Penrose since 1964 (TSP and Pb) and MfE’s monitoring at this site is scheduled to be decommissioned in July 2011. Previously, there were monitoring sites in Mt Eden and Kingsland, both of which are older residential areas just south of the city centre. Monitoring at Mt Eden was conducted between 1982 and 2004. In October 2004 the Mt Eden site was decommissioned pending redevelopment of the land and replaced by the Kingsland site at Kowhai Intermediate. In September 2007, the Kowhai, Kingsland site was decommissioned due to the construction of a swimming pool complex.
Christchurch monitoring from 1989 to 2002 monitoring was undertaken at a site located in the older residential area of St Albans which is just north of the city centre. The monitoring station was relocated in November 2002 due to impending redevelopment of the site. The current site is in Burnside / Bishopdale which is a newer residential area to the north-west of the city centre. This Burnside air quality monitoring site was decommissioned in January 2011.
Environment Canterbury provided sampling services for MfE for gravimetric and passive monitoring methods at the Christchurch sites (Coles Place, St Albans and Greers Road Burnside). This type of monitoring also ceased in January 2011.

This report includes graphical and statistical presentations of the data as well as any data collection issues that may have arisen during the monitoring period.
All data in this report has been validated. Quality assurance checks have been carried out to ensure that invalid and calibration data is not reported.
3
Air Pollutants Monitored

3.1
Carbon monoxide

Carbon monoxide (CO) is a colourless, odourless, toxic gas formed as a product of incomplete combustion in the burning of fossil fuels. The main sources of carbon monoxide in most urban areas are vehicle exhaust emissions and home heating.
 Elevated concentrations are mainly found in areas of significant traffic congestion, particularly at busy intersections on inner-city streets.
Carbon monoxide inhibits oxygen uptake in the blood. Prolonged exposure at moderate levels can lead to symptoms such as headaches and dizziness, while at high levels it can lead to loss of consciousness and even death.
3.2
Nitrogen oxides

Nitrogen oxides incorporate several substances that exist in the atmosphere, which are collectively referred to as NOx. The two main oxides are nitrogen dioxide (NO2) and nitric oxide (NO). NO is less toxic than NO2 but may oxidise to NO2 in the atmosphere.
Nitrogen oxides are formed in most combustion processes by oxidation of the nitrogen present in the atmosphere. Nitric oxide is the predominant primary product but, as indicated, this can then be oxidised to nitrogen dioxide in ambient air. Emissions from motor vehicles are the major source of NOx in most parts of the country, although power stations and other large combustion units may be significant localised sources as well.

NOx is also an important air pollutant because of its role in photochemical smog. NO2 is a reddish brown gas, and has synergistic health effects with other pollutants such as SO2 and particulate matter.

The main health effects of the oxides of nitrogen are due to NO2, which is a respiratory irritant. Nitric oxide is believed to be quite harmless at the levels normally encountered in urban air.
3.3
Sulfur dioxide

Sulfur dioxide is an acidic gas with a pungent odour which is mainly produced by the burning of fossil fuels. The gas is quite corrosive and can cause damage to buildings and other materials.
It can also have significant effects on the human respiratory system. Inhalation of high ambient concentrations of sulfur dioxide can cause stimulation of the nerves in the air passages, resulting in a reflex cough, irritation and chest tightness.
In addition, sulfur dioxide can also cause narrowing of the air passages, particularly in people suffering from asthma and chronic lung disease. These people frequently have narrowed airways and any further restriction will have a more significant effect compared to people with uncompromised respiratory systems.
3.4
Volatile organic compounds

Volatile organic compounds (VOCs) are chemicals that easily evaporate at room temperature. The term ‘organic’ indicates that the compounds contain carbon. VOCs include a wide range of carbon-based molecules, such as some aldehydes, ketones and hydrocarbons.
The Ministry for the Environment has compiled a list of priority contaminants, based on a review of international literature. The priority list includes the VOCs benzene and 1,3-butadiene and provides ambient air quality guidelines for these contaminants.
 VOCs can produce neurological impairment, and benzene is well established as a carcinogen.

3.5
Particulate matter

Particulate matter (PM) can be a significant air pollutant that is associated with a variety of health and environmental effects.

Sources of particulates vary widely from location to location reflecting the wide range of emission sources that contribute to particulate concentrations in New Zealand. Typical sources of various particulate sizes can include:

· Fine particulates (less than 2.5 μm) emitted as a result of incomplete fuel combustion from road vehicles, power generation, industrial processes and domestic heating appliances. Fine particulates also form by chemical reactions in the atmosphere. These comprise largely of sulfates and nitrates.

· Coarse particulates (between 2.5 and 10 μm) that arise from a wide range of sources, including re-suspended dust from road vehicles, construction works, mineral extraction processes, wind-blown dust and soils, sea salt and biological particulates such as pollen and fungal spores.

Particulate matter refers to numerous substances that exist in the atmosphere. It is a somewhat complex pollutant, encompassing a wide range of chemically and physically diverse substances. Particulate matter includes all solid and liquid aerosols that exist in ambient conditions.
As part of the GEMS/AMIS programme two particle size fractions are monitored:

· particulate matter less than 10 μm (PM10)
· total suspended particulates (TSP).

3.5.1
Particulate matter as PM10
The main effect of PM10 is on human health as PM10 can be inhaled into the respiratory system. PM10 can become deposited in the trachea bronchial region, where asthma attacks are triggered.
3.5.2
Total suspended particulate
TSP consists of all particles which range in size up to 100 µm in diameter. TSP is sufficiently small enough to be inhaled; however, the larger particles (10–100 µm) are readily filtered out in the nasal cavity. TSP has an effect on both the aesthetic and health qualities of ambient air.
3.6
Lead

Lead is a toxic metal present in the atmosphere in its elemental form and one of the principal sources has been motor vehicle emissions. Historically, lead was included in petrol as a catalyst for combustion, but has been removed from fuel supplies since 1996.
As a result, atmospheric concentrations of lead have dropped markedly since this time. To reflect this, from October 2000 monitoring of lead was reduced from monthly samples to samples taken over a 3-month period during winter (June–August) only.
Health effects of lead in air are caused when the small particulates are inhaled or swallowed. It can harm the central nervous system, kidneys, and blood cells. Growing children and foetuses are more at risk resulting in possible physical and mental problems.
4
Ambient Air Quality Guidelines and Standards

In October 2004, the Ministry for the Environment introduced the National Environmental Standards (NES) for Air Quality. The NES includes five standards for ambient (outdoor) air quality. Four of these standards and some New Zealand guidelines are described in Table 1 below. The fifth standard is not included as it relates to ozone, which is not measured at the sites discussed in this report.

The criteria used for TSP, 60 µg/m3 (seven-day average), was previously used by the Ministry of Health (MoH). This has been superseded by the ambient air quality guidelines but is still useful for analysing the results of the monitoring data.
Table 1:
National environmental standards and guidelines

	Air pollutant
	Averaging period
	National Environmental Standards 2004
	Ministry for the Environment Ambient Air Quality Guidelines 2002 and other
	National Environmental Standards permissible excess

	Carbon monoxide
	8-hour average
	10 mg/m3
	10 mg/m3
	One 8-hour period in a 12‑month period

	
	1-hour average
	
	30 mg/m3
	

	Nitrogen dioxide
	24-hour average
	
	100 µg/m3
	

	
	1-hour average
	200 µg/m3
	200 µg/m3
	9 hours in a 12-month period

	Sulfur dioxide
	24-hour average
	
	120 µg/m3
	

	
	1-hour average
	350 µg/m3
	350 µg/m3
	9 hours in a 12-month period

	
	1-hour average
	570 µg/m3
	
	Not to be exceeded at any time

	Benzene 
	
	
	
	

	– Year 2002
	Annual average
	
	10 µg/m3
	

	– Year 2010
	Annual average
	
	3.6 µg/m3
	

	1,3-Butadiene
	Annual average
	
	2.4 µg/m3
	

	Fine particulate (PM10)
	Annual average
	
	20 µg/m3
	

	
	24-hour average
	50 µg/m3
	50 µg/m3
	One 24-hour period in a 12‑month period

	Total suspended particulate (TSP)
	7-day average
	
	60 µg/m3 (MoH)
	

	Lead
	3-month average
	
	0.2 µg/m3
	


5
Monitoring Sites

5.1
Site descriptions

A brief description of all the monitoring sites in the GEMS/AMIS air quality monitoring programme is given below. This includes one Auckland site at Gavin Street in Penrose and two sites in Christchurch at Greers Road in Burnside and Coles Place in St Albans. (Note at Coles Place, only TSP, Lead and VOCs are monitored by MfE).

5.1.1
MfE Gavin Street, Penrose, Auckland – Site AKL009
	Site name
	MfE Gavin Street, Penrose
	Site ID
	AKL009

	Address
	Transpower, Gavin Street, Penrose, Auckland
	Site class
	Industrial – dense 
Traffic – peak

	Description

	This site is operated by Watercare and is located within the grounds of the Transpower NZ Ltd electrical substation on Gavin Street, Penrose. It is representative of road, vehicle and industrial emissions in the Penrose area which lies to the southeast of Auckland City and is also approximately 120 metres northeast of the Southern Motorway. There are residential properties immediately to the northeast of the site. During 2003, parallel monitoring was undertaken between this site and the neighbouring ACI site on Great South Road, in Penrose with a view to consolidating all monitoring at the Gavin Street site early in 2004. Other pollutants measured at this site by Auckland Council will continue to be measured.

	Pollutants monitored by MfE
	CO
	NO2
	SO2
	VOCs

	
	N
	Y
	Y
	Y

	
	PM10
	TSP
	Lead
	

	
	Y
	Y
	Y
	

	Meteorological parameters monitored
	Wind speed
	Wind direction
	Relative humidity
	

	
	Y
	Y
	Y
	

	
	Temperature (6m)
	Temperature (10m)
	Temperature (1.5m)
	

	
	Y
	N
	N
	

	Location map

[image: image1.jpg]*’Ea}SECare

services l[imited



[image: image66.png][image: image67.jpg]New Zealand Government




[image: image3.jpg]om 150 200 7 300 o





	Photograph

[image: image4.jpg]





5.1.2
MfE Greers Road, Burnside, Christchurch – Site CAN002

	Site name
	MfE Greers Road, Burnside
	Site ID
	CAN002

	Address
	Transpower, Greers Road, Burnside, Christchurch
	Site class
	Residential – neighbourhood

	Description

	This site is operated by Watercare and is located in a paddock to the rear of the Transpower NZ Ltd electrical substation on Greers Road, and is surrounded by residential properties on four sides. Greers Road is approximately 100 metres to the north west of the site. It is representative of emissions arising from domestic properties in the newer suburban areas of Burnside and Bishopdale which lie to the northwest of Christchurch city centre. The site was commissioned in November 2002 and replaces the former GEMS/AMIS site which was located off Madras Street, St Albans. The site was decommissioned in January 2011.

	Pollutants monitored by MfE
	CO
	NO2
	SO2
	VOCs

	
	Y
	Y
	Y
	Y

	
	PM10
	TSP
	Lead
	

	
	Y
	N
	N
	

	Meteorological parameters monitored
	Wind speed
	Wind direction
	Relative humidity
	

	
	Y
	Y
	Y
	

	
	Temperature (6m)
	Temperature (10m)
	Temperature (1.5m)
	

	
	N
	Y
	Y
	

	Location map

[image: image5.wmf]MfE Greers 

Road, Burnside


	Photograph

[image: image6.jpg]





5.1.3
MfE Coles Place, St Albans, Christchurch – Site CAN003

	Site name
	MfE Coles Place, St Albans
	Site ID
	CAN003

	Address
	Coles Place, St Albans, Christchurch
	Site class
	Residential – neighbourhood

	Description

	This site is operated by Environment Canterbury and is located on an area of public open space at the end of Coles Place. It is surrounded by residential properties on four sides. It is representative of emissions arising from domestic properties in the older suburban area of St Albans which lies to the north of Christchurch city centre. As of January 2011 the MfE ceased VOC, TSP and lead monitoring. Other pollutants are continued to be monitored at this site by Environment Canterbury. 

	Pollutants monitored by MfE
	CO
	NO2
	SO2
	VOCs

	
	N
	N
	N
	Y

	
	PM10
	TSP
	Lead
	

	
	N
	Y
	Y
	

	Meteorological parameters monitored
	Wind speed
	Wind direction
	Relative humidity
	

	
	N
	N
	N
	

	
	Temperature (6m)
	Temperature (10m)
	Temperature (1.5m)
	

	
	N
	N
	N
	

	Location map

[image: image7.wmf]Coles Place,

St Albans


	Photograph

[image: image8.wmf]


5.2
Air pollutants monitored under the GEMS/AMIS programme
An overview of the air pollutants monitored at each site during 2010, as part of the GEMS/AMIS programme, is given in Table 2 below.

Table 2:
Air pollutants monitored under the GEMS/AMIS programme 2010
	Site
	CO
	NO2
	SO2
	VOC
	PM10
	TSP
	Lead*

	Transpower, Gavin Street, Penrose, Auckland
AKL009
	
	ü
	ü
	ü
	ü
	ü
	ü

	Greers Road, Burnside, Christchurch
CAN002
	ü
	ü
	ü
	ü
	ü
	
	

	Coles Place, St Albans, Christchurch
CAN003
	
	
	
	ü
	
	ü
	ü


Note:
*
Lead is measured from the TSP filters collected during the months of June, July and August only.

6
Methods

6.1
Quality assurance

All monitoring services are undertaken by Watercare accordance with the MfE ‘Good Practice Guide for Air Quality Monitoring and Data Management’ (or GPG) and wherever applicable the appropriate Australian/New Zealand and US EPA monitoring methods.

Watercare is accredited by International Accreditation New Zealand (IANZ) for the following methods:

· AS 3580.7.1-1992 ‘Method 7.1: Determination of carbon monoxide – direct-reading instrumental method’
· AS 3580.5.1-1993 ‘Method 5.1: Determination of oxides of nitrogen – chemiluminescence method’
· AS 3580.4.1-2008 ‘Method 4.1: Determination of sulfur dioxide – direct reading instrumental method’
· AS/NZS 3580.9.3-2003 ‘Method 9.3: Determination of ambient particulates (gravimetric method) – TSP high volume sampling’
· AS/NZS 3580.9.6-2003 ‘Method 9.6: Determination of ambient particulates (gravimetric method) – PM10 high volume sampling’
· AS 3580.9.11 – 2008 ‘Method 9.11: Determination of suspended particulate matter – PM10 beta attenuation monitors’
· APHA Method 3030 for the lead digest

· US EPA Method 200.8 (modified) for lead analysis by ICPMS.
As part of the GEMS/AMIS programme, Watercare provides the following monitoring services:

· instrument operation, calibration and maintenance. This includes the use of automatic daily calibration systems for all continuous ambient gas monitors ensuring that the requirements of the relevant Australian Standards for weekly calibration of continuous analysers are met

· site maintenance as well as commissioning new sites and decommissioning old sites, when necessary
· data logging, polling, checking, re-scaling, validation, ratification and reporting. This encompasses the entire data quality assurance process ensuring that the final data set reported is fit for the purpose of the GEMS/AMIS programme.

6.2
Analytical methods

6.2.1
Carbon monoxide

Measurements are made in accordance with AS 3580.7.1-1992 ‘Determination of Carbon Monoxide – Direct Reading Instrumental Method’. The performance of the instrument is checked using an automatic calibration system ensuring compliance with the method which requires instrumentation to be calibrated on a weekly basis.

The instrument itself is an infrared absorption gas analyser which continuously measures carbon monoxide. This allows data to be analysed and reported over a variety of average periods, including 10 minute, 24 hour and one year.

6.2.2
Nitrogen oxides

Measurements are made in accordance with AS 3580.5.1-1993 ‘Determination of Oxides of Nitrogen – Chemiluminescence Method’. The performance of the instrument is checked using an automatic calibration system ensuring compliance with the method which requires instrumentation to be calibrated on a weekly basis.

The instruments themselves are chemiluminescence gas analysers which continuously measure nitrogen oxides. This allows data to be analysed and reported over a variety of average periods, including 10 minute, 24 hour and one year.

6.2.3
Sulfur dioxide
Measurements are made in accordance with AS 3580.4.1-2008 ‘Determination of Sulfur Dioxide – Direct Reading Instrumental Method’. The performance of the instrument is checked using an automatic calibration system ensuring compliance with the method which requires instrumentation to be calibrated on a weekly basis.

The instruments themselves are UV fluorescence gas analysers which continuously measure sulfur dioxide. This allows data to be analysed and reported over a variety of average periods, including 10 minute, 24 hour and one year.

6.2.4
Volatile organic compounds (VOCs)
VOCs are measured using passive 3M organic vapour monitor badges, which are exposed for three-month periods (January–March, April–June, July–September and October–December). The VOCs diffuse on to the badges, which are coated with activated carbon. Following exposure, the samples are forwarded to AsureQuality who extract the VOCs using carbon disulfide and analyse them using gas chromatography-mass spectrometry (GC-MS). AsureQuality is IANZ accredited for VOCs by GC-MS using National Institute for Occupational Safety and Health (NIOSH) Method 1500.
The subsequent results are calculated in accordance with the manufacturer’s instructions (3M Technical Data Bulletin 1028) allowing results to be converted to standard temperature (0 °C) and pressure (1 atm).
Note: Investigations have determined that samples of 1,3-butadiene are unstable when held above -4.41 (C with significant reverse desorption occurring. Due to the potential for error over a three-month exposure period, 1,3-butadiene has not been analysed and reported. Alternative methods of measuring 1,3-butadiene are OSHA Method 56 and NIOSH Method 1024. Both of these involve the use of solid sorbent coconut charcoal tubes.

6.2.5
Particulate matter as PM10
Measurements are made in accordance with AS 3580.9.11–2008 ‘Determination of suspended particulate matter – PM10 beta attenuation monitors’ (BAM). The instrument is fitted with a size-selective PM10 head with an inlet temperature of 40 °C. The BAM continuously measures the intensity of beta particles passing through a filter tape. This allows PM10 data to be analysed and reported over a variety of average periods, including 10 minute, 24 hour and one year.
6.2.6
Total suspended particulate matter (TSP)

TSP monitoring is in accordance with Watercare’s Air Quality Group Test Method T101. The method was based on Australian Standard 2724.2 – 1987 and British Standard 1747 and was developed by the Ministry of Health (MoH). It is a gravimetric method of measuring particulates and is modelled on the system for high volume sampling as described in AS/NZS 3580.9.3-2003.
The equipment used to collect TSP involves ambient air being pulled through a 55 mm glass fibre filter by a vacuum pump. A gas meter is attached and is used to measure the air volume (m³) drawn through the filter. The filter is held in a cassette and is covered with a rain shelter at 3.5 m above ground. Preweighed filters are collected and sampled on a weekly or 7 day cycle every Thursday.
This method is different to that recommended in the Ministry for the Environment’s ‘Good Practice Guide for Air Quality Monitoring and Data Management 2010’. The method has been used to measure TSP since 1964 and the results will vary when compared to more modern techniques.
6.2.7
Lead

Lead is sampled during the winter months, June–August, using the same samples acquired to measure TSP. Analysis of lead is performed by Watercare Laboratory Services according to APHA Method 3030 and US EPA Method 200.8 modified. This involves analysing each individual TSP filter exposed during the winter period using mixed acid digestion. This sample is then analysed for lead using ICP-MS. The concentration of lead is then determined from the amount of lead detected and the total volume of air sampled during that sample period. Concentrations are averaged for the 3-month monitoring period.
7
Results and Discussion

7.1
Site performance and quality assurance

Monthly and annual site performance and explanations are shown in Table 3, based on 10 minute averages for continuously monitored data. Per cent of valid data (V) is defined as the per cent of valid data following quality assurance adjustments. Per cent of captured data (C) is the per cent of valid data excluding calibration and maintenance.
The 2000 MfE Good Practice Guide for Air Quality Monitoring and Data Management suggests that it is difficult to reach anything close to 100% valid data for long-term monitoring. As such site performance has been evaluated against a target of 95% for capture data and 75% for valid data.

The performance of continuously monitored pollutant instruments during 2010 was generally very good. All sites had annual valid data greater than 75% and data capture rates greater than 95% with the exception of one VOC sample at Burnside.
Table 3:
Percentage valid and capture data 2010
	Analyte
	Site
	Percentage valid data (V) and percentage data capture (C)

	
	
	Jan
	Feb
	Mar
	Apr
	May
	Jun
	Jul
	Aug
	Sep
	Oct
	Nov
	Dec
	Annual

	
	
	V
	C
	V
	C
	V
	C
	V
	C
	V
	C
	V
	C
	V
	C
	V
	C
	V
	C
	V
	C
	V
	C
	V
	C
	V
	C

	CO
	Burnside
	98.5
	100
	97.2
	99.1
	98.5
	100
	98.4
	100
	98.1
	100
	98.5
	100
	98.2
	99.9
	89.9
	91.8d
	97.2
	98.7
	97.6
	100
	97.8
	99.5
	97.7
	99.2
	97.3
	99.0

	NO2
	Penrose
	97.6
	99.7
	98.0
	99.8
	97.8
	100
	98.0
	100
	98.1
	100
	97.6
	100
	98.1
	100
	98.1
	100
	97.4
	100
	97.5
	100
	79.8
	81.6f
	97.8
	100
	96.3
	98.4

	NO2
	Burnside
	98.2
	100
	97.8
	99.8
	98.0
	100
	97.9
	100
	97.9
	100
	98.2
	100
	97.4
	100
	89.6
	91.7d
	87.5
	89.0 e
	97.1
	100
	81.1
	83.1g
	97.6
	100
	94.9
	97.0

	SO2
	Penrose
	97.1
	99.8
	97.4
	99.8
	97.9
	100
	97.6
	99.8
	97.9
	99.9
	97.9
	100
	98.1
	100
	98.2
	100
	98.1
	100
	97.8
	100
	97.6
	99.9
	97.1
	99.5
	97.7
	99.9

	SO2
	Burnside
	97.3
	100
	97.5
	99.8
	97.7
	100
	97.8
	100
	97.7
	100
	97.8
	100
	97.7
	100
	89.5
	91.8d
	96.7
	99.0
	97.3
	100
	97.6
	100
	97.6
	100
	96.8
	99.2

	PM10
	Penrose
	98.5
	99.0
	98.6
	99.3
	98.9
	99.6
	98.8
	99.8
	99.2
	99.9
	99.3
	100
	83.9
	85.0 c
	99.0
	99.7
	99.2
	99.9
	96.6
	97.2
	99.0
	99.7
	98.8
	99.5
	97.4
	97.6

	PM10
	Burnside
	98.4
	99.1
	98.9
	99.6
	98.4
	99.1
	96.1
	97.6
	98.5
	99.7
	98.9
	99.6
	99.0
	99.7
	89.9
	90.5d
	97.1
	98.0
	98.2
	99.4
	98.6
	99.3
	98.6
	99.3
	97.5
	97.8

	VOC
	Penrose
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100

	VOC
	Burnside
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	0h
	0h
	75.0
	75.0

	VOC
	St Albans
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100

	TSP
	Penrose
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	75.0
	75.0b
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	98.1
	98.1

	TSP
	St Albans
	75.0
	75.0a
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	100
	75.0
	75.0a
	100
	100
	100
	100
	75.0
	75.0a
	94.2
	94.2

	Lead
	Penrose
	–
	–
	–
	–
	–
	–
	–
	–
	–
	–
	75.0
	–
	100
	–
	100
	–
	–
	–
	–
	–
	–
	–
	–
	–
	92.3
	–

	Lead
	St Albans
	–
	–
	–
	–
	–
	–
	–
	–
	–
	–
	100
	–
	100
	–
	100
	–
	–
	–
	–
	–
	–
	–
	–
	–
	100
	–


a
TSP – Sampling error
b
TSP – No power
c
PM10 – Pump faulty

d
CO, NO2, SO2, and PM10 – Power connection failure

e
NO2 – Christchurch earthquake no data
f
NO2 – Moly converter faulty
g
NO2 – Ozone flow faulty
h
VOC – Sample contaminated

7.2
Carbon monoxide (CO) 2010
CO was monitored at Greers Road, Burnside. One hour and 8-hour rolling averages have been calculated from 10-minute averages recorded by the instrument.

Summary statistics for CO and their dates are described below.

	Site
	Maximum
1-hour average
(mg/m3)
	99.9 percentile
1-hour average
(mg/m3)
	Maximum
8-hour average
(mg/m3)
	99.9 percentile
8-hour average
(mg/m3)

	Greers Road, Burnside
	8.4
(21 Jun 23:00)
	6.3
	5.8
(22 Jun 02:00)
	4.7


At Greers Road, Burnside during 2010 there were no exceedences of the ambient air quality 1‑hour guideline (30 mg/m3) or the 8-hour national environmental standard (10 mg/m3). The 2010 CO results are shown in Figures 7 to 11.

7.3
Nitrogen oxides (NO2 and NO) 2010
Oxides of nitrogen were monitored at Gavin Street, Penrose and Greers Road, Burnside. One hour and 24-hour averages have been calculated from 10-minute averages recorded by the instruments.
Summary statistics for NO2 and their dates for each site are described below.

	Site
	Maximum
1-hour average
(µg/m3)
	99.9 percentile
1-hour average
(µg/m3)
	Maximum
24-hour average
(µg/m3)
	99.5 percentile
24-hour average
(µg/m3)

	Gavin Street, Penrose
	92.9
(16 Jun 09:00)
	75.5
	54.5
(16 Jun)
	52.6

	Greers Road, Burnside
	58.1
(14 Jun 11:00)
	53.3
	30.7
(14 May)
	28.3


There were no exceedences of the NO2 ambient air quality 1-hour standard (200 µg/m3) or the 24-hour guideline (100 µg/m3) during 2010 at either site.
The 2010 NOx results for Penrose are shown in Figures 12 to 21 and results for Burnside are shown in Figures 22 to 31.

7.4
Sulfur dioxide (SO2) 2010
Sulfur dioxide was monitored at Gavin Street, Penrose and Greers Road, Burnside. One hour and 24-hour averages have been calculated from 10-minute averages recorded by the instruments.
Summary statistics for SO2 and their dates for each site are described below.

	Site
	Maximum
1-hour average
(µg/m3)
	99.9 percentile
1-hour average
(µg/m3)
	Maximum
24-hour average
(µg/m3)
	99.5 percentile
24-hour average
(µg/m3)

	Gavin Street, Penrose
	59.4
(13 Oct 05:00)
	38.5
	27.7
(13 Oct)
	16.2

	Greers Road, Burnside
	22.4
(28 Jul 23:00)
	18.4
	8.7
(24 Jun)
	7.4


There were no exceedences of the SO2 national environmental 1-hour standard (350 µg/m3) or the 24-hour guideline (120 µg/m3) during 2010 at either site.

The 2010 SO2 results for Penrose are shown in Figures 32 to 36 and results for Burnside are shown in Figures 37 to 41.
7.5
Volatile organic compounds (VOC) January–December 2010
Monitoring of VOCs was conducted at three sites; Gavin Street, Penrose; Greers Road, Burnside; and Coles Place, St Albans. VOC monitoring utilises passive sampling badges exposed over a 3-month period. Results for each 2010 quarter are shown in Tables 4 to 7.
The benzene guideline 2002 was 10 µg/m3 as an annual average, with an average value of 3.6 µg/m3 to be achieved by 2010. The 2010 6-month and 12-month averages are described below. The benzene annual averages from all the sites are less than the annual average guideline (3.6 µg/m3) and are graphed in Figure 1.

	Site
	Six-month average (January–June 2010) benzene
(µg/m3)
	Six-month average (July–December 2010)
benzene
(µg/m3)
	2010 annual average

benzene
(µg/m3)

	Gavin Street, Penrose
	0.9
	0.7
	0.8

	Greers Road, Burnside
	1.0
	1.0
	1.0

	Coles Place, St Albans
	1.1
	1.0
	1.0


Table 4:
VOC results (January–March 2010)

	Jan–Feb–Mar 2010
	Limit of detection
	Results (µg/m3)

	Analyte
	(µg/m3)
	Coles Place
	Burnside
	Penrose

	Target VOCs
	
	
	
	

	ethanol
	ND
	
	
	

	isopropyl alcohol
	ND
	
	
	

	acetone
	ND
	
	
	

	pentane
	2.6
	ND
	ND
	3.2

	dichloromethane
	ND
	
	
	

	butan-2-one
	ND
	
	
	

	hexane
	ND
	
	
	

	ethyl acetate
	ND
	
	
	

	trichloromethane
	ND
	
	
	

	1,1,1-trichloroethane
	ND
	
	
	

	n-butanol
	ND
	
	
	

	benzene
	0.3
	0.4
	0.4
	0.5

	2-methylhexane
	ND
	
	
	

	2,3-dimethylpentane
	ND
	
	
	

	3-methylhexane
	ND
	
	
	

	heptane
	ND
	
	
	

	trichloroethene
	ND
	
	
	

	propyl acetate
	ND
	
	
	

	methylcyclohexane
	ND
	
	
	

	4-methylpentan-2-one
	ND
	
	
	

	toluene
	0.3
	2.2
	1.8
	4.0

	octane
	ND
	
	
	

	tetrachloroethene
	ND
	
	
	

	butyl acetate
	ND
	
	
	

	ethylbenzene
	0.3
	ND
	ND
	0.6

	m+p-xylene
	0.3
	1.1
	0.9
	2.0

	styrene
	ND
	
	
	

	o-xylene
	0.3
	0.4
	ND
	0.7

	nonane
	ND
	
	
	

	alpha pinene
	ND
	
	
	

	propylbenzene
	ND
	
	
	

	1,3,5-trimethylbenzene
	ND
	
	
	

	beta pinene
	ND
	
	
	

	decane
	ND
	
	
	

	1,2,4-trimethylbenzene
	0.7
	ND
	ND
	0.7

	limonene
	ND
	
	
	

	undecane
	ND
	
	
	

	dodecane
	ND
	
	
	

	tetradecane
	ND
	
	
	


ND = Not detected.
Table 5:
VOC results (April–June 2010)

	Apr–May–Jun 2010
	Limit of detection
	Results (µg/m3)

	Analyte
	(µg/m3)
	Coles Place
	Burnside
	Penrose

	Target VOCs
	
	
	
	

	ethanol
	ND
	
	
	

	isopropyl alcohol
	ND
	
	
	

	acetone
	ND
	
	
	

	pentane
	ND
	
	
	

	dichloromethane
	ND
	
	
	

	butan-2-one
	ND
	
	
	

	hexane
	0.5
	1.1
	0.8
	1.3

	ethyl acetate
	ND
	
	
	

	trichloromethane
	ND
	
	
	

	1,1,1-trichloroethane
	ND
	
	
	

	n-butanol
	ND
	
	
	

	benzene
	0.2
	1.9
	1.6
	1.4

	2-methylhexane
	0.5
	0.6
	ND
	ND

	2,3-dimethylpentane
	ND
	
	
	

	3-methylhexane
	0.5
	0.7
	ND
	0.6

	heptane
	0.5
	0.7
	ND
	0.5

	trichloroethene
	ND
	
	
	

	propyl acetate
	ND
	
	
	

	methylcyclohexane
	ND
	
	
	

	4-methylpentan-2-one
	ND
	
	
	

	toluene
	0.3
	8.0
	5.8
	8.3

	octane
	ND
	
	
	

	tetrachloroethene
	ND
	
	
	

	butyl acetate
	ND
	
	
	

	ethylbenzene
	0.3
	1.2
	0.9
	1.3

	m+p-xylene
	0.3
	3.9
	2.9
	4.6

	styrene
	ND
	
	
	

	o-xylene
	0.3
	1.4
	1.1
	1.6

	nonane
	ND
	
	
	

	alpha pinene
	ND
	
	
	

	propylbenzene
	ND
	
	
	

	1,3,5-trimethylbenzene
	ND
	
	
	

	beta pinene
	0.7
	ND
	0.9
	ND

	decane
	ND
	
	
	

	1,2,4-trimethylbenzene
	0.6
	1.8
	1.4
	1.7

	limonene
	ND
	
	
	

	undecane
	ND
	
	
	

	dodecane
	ND
	
	
	

	tetradecane
	ND
	
	
	


ND = Not detected.
Table 6:
VOC results (July–September 2010)

	Jul–Aug–Sept 2010
	Limit of detection
	Results (µg/m3)

	Analyte
	(µg/m3)
	Coles Place
	Burnside
	Penrose

	Target VOCs
	
	
	
	

	ethanol
	ND
	
	
	

	isopropyl alcohol
	ND
	
	
	

	acetone
	ND
	
	
	

	pentane
	2.3
	3.6
	4.6
	7.6

	dichloromethane
	ND
	
	
	

	butan-2-one
	ND
	
	
	

	hexane
	0.5
	1.0
	1.5
	0.9

	ethyl acetate
	ND
	
	
	

	trichloromethane
	ND
	
	
	

	1,1,1-trichloroethane
	ND
	
	
	

	n-butanol
	ND
	
	
	

	benzene
	0.2
	2.0
	2.0
	1.2

	2-methylhexane
	ND
	
	
	

	2,3-dimethylpentane
	ND
	
	
	

	3-methylhexane
	0.5
	0.6
	0.6
	ND

	heptane
	0.5
	ND
	0.7
	ND

	trichloroethene
	ND
	
	
	

	propyl acetate
	ND
	
	
	

	methylcyclohexane
	ND
	
	
	

	4-methylpentan-2-one
	ND
	
	
	

	toluene
	0.3
	6.5
	6.0
	5.8

	octane
	ND
	
	
	

	tetrachloroethene
	ND
	
	
	

	butyl acetate
	ND
	
	
	

	ethylbenzene
	0.3
	0.9
	0.8
	1.0

	m+p-xylene
	0.3
	2.7
	2.4
	3.3

	styrene
	ND
	
	
	

	o-xylene
	0.3
	1.0
	0.9
	1.1

	nonane
	ND
	
	
	

	alpha pinene
	ND
	
	
	

	propylbenzene
	ND
	
	
	

	1,3,5-trimethylbenzene
	ND
	
	
	

	beta pinene
	0.7
	ND
	0.8
	ND

	decane
	ND
	
	
	

	1,2,4-trimethylbenzene
	0.6
	1.5
	1.3
	1.1

	limonene
	ND
	
	
	

	undecane
	ND
	
	
	

	dodecane
	ND
	
	
	

	tetradecane
	ND
	
	
	


ND = Not detected.
Table 7:
VOC results (October–December 2010)

	Oct–Nov–Dec 2010
	Limit of detection
	Results (µg/m3)

	Analyte
	(µg/m3)
	Coles Place
	Burnside1
	Penrose

	Target VOCs
	
	
	
	

	ethanol
	ND
	
	
	

	isopropyl alcohol
	ND
	
	
	

	acetone
	ND
	
	
	

	pentane
	2.1
	ND
	
	2.8

	dichloromethane
	ND
	
	
	

	butan-2-one
	ND
	
	
	

	hexane
	0.4
	ND
	
	0.5

	ethyl acetate
	ND
	
	
	

	trichloromethane
	ND
	
	
	

	1,1,1-trichloroethane
	ND
	
	
	

	n-butanol
	ND
	
	
	

	benzene
	0.2
	0.4
	
	0.5

	2-methylhexane
	ND
	
	
	

	2,3-dimethylpentane
	ND
	
	
	

	3-methylhexane
	ND
	
	
	

	heptane
	ND
	
	
	

	trichloroethene
	ND
	
	
	

	propyl acetate
	ND
	
	
	

	methylcyclohexane
	ND
	
	
	

	4-methylpentan-2-one
	ND
	
	
	

	toluene
	0.2
	2.4
	
	3.8

	octane
	ND
	
	
	

	tetrachloroethene
	ND
	
	
	

	butyl acetate
	ND
	
	
	

	ethylbenzene
	0.3
	ND
	
	0.5

	m+p-xylene
	0.3
	0.6
	
	1.6

	styrene
	ND
	
	
	

	o-xylene
	0.3
	ND
	
	0.5

	nonane
	ND
	
	
	

	alpha pinene
	ND
	
	
	

	propylbenzene
	ND
	
	
	

	1,3,5-trimethylbenzene
	ND
	
	
	

	beta pinene
	ND
	
	
	

	decane
	ND
	
	
	

	1,2,4-trimethylbenzene
	ND
	
	
	

	limonene
	ND
	
	
	

	undecane
	ND
	
	
	

	dodecane
	ND
	
	
	

	tetradecane
	ND
	
	
	


1
Sample contaminated
ND = Not detected.
Figure 1:
MfE Benzene annual average 2010
[image: image9.emf]0

2

4

6

8

10

12

Coles Place Burnside Penrose

Site

Benzene µg/m³

AAQG 2010 = 3.6 µg/m³

AAQG 2002 = 10.0 µg/m³


7.6
Particulate matter (PM10) 2010
PM10 was monitored at Gavin Street, Penrose and Greers Road, Burnside. Twenty-four-hour averages have been calculated from 10-minute averages recorded by the instruments.
Summary statistics for PM10 and their dates for each site are described below.

	Site
	Maximum 24-hour average
(µg/m3)
	99.5 percentile 24-hour average
(µg/m3)

	Gavin Street, Penrose
	43
(22 Nov)
	36

	Greers Road, Burnside
	70
(28 Jul)
	61


There were no exceedences of the 24-hour national environmental standard (50 µg/m3) at the Auckland site during 2010. The 2010 PM10 results for Penrose are shown in Figures 42 to 44.
At Greers Road, Burnside, there were eight exceedences of the 24-hour standard. Each exceedence and the date of the exceedence are listed in Table 8 below. As there were multiple exceedences at the Christchurch site, more data analysis was carried out and reported in Section 7.9. The 2010 PM10 results for Burnside are shown in Figures 45 to 47.
All exceedences occurred over the winter period (June to August), a time when wood burning is widely used for domestic heating. Cold winter conditions strongly influence air pollution in the region especially in calm conditions.
Table 8:
Greers Road, Burnside PM10 exceedences of the daily NES 2010
	Date
	Burnside PM10
(µg/m3)

	21-Jun-10
	61

	1-Jul-10
	59

	3-Jul-10
	50

	11-Jul-10
	58

	13-Jul-10
	56

	27-Jul-10
	61

	28-Jul-10
	70

	29-Jul-10
	56


Note: National environmental standard for PM10 = 50 µg/m³

7.7
Total suspended particulates (TSP) 2010
TSP is measured as a 7-day average at Gavin Street, Penrose and Coles Place, St Albans. Maximum results and their dates (7-day ending period) for each site are described below.

	Site
	Maximum 7-day average
(µg/m3)

	Gavin Street, Penrose
	26
(10 Mar)

	Coles Place, St Albans
	44
(08 Jul)


There were no exceedences of the MoH guideline of 60 µg/m3 at either site. The TSP concentrations in Auckland are shown in Figures 2 and 3 while Christchurch TSP concentrations are shown in Figures 4 and 5.

Figure 2:
Auckland TSP 7-day average January–December 2010
[image: image10.emf]0

10

20

30

40

50

60

70

80

90

100

Jan-10 Mar-10 May-10 Jul-10 Sep-10 Nov-10

Date

TSP µg/m³

Gavin Street, Penrose

MoH Guideline = 60µg/m³


Figure 3:
Auckland TSP 7-day average 1994–2010

[image: image11.emf]0

10

20

30

40

50

60

70

80

90

100

Jun-

94

Jun-

95

Jun-

96

Jun-

97

Jun-

98

Jun-

99

Jun-

00

Jun-

01

Jun-

02

Jun-

03

Jun-

04

Jun-

05

Jun-

06

Jun-

07

Jun-

08

Jun-

09

Jun-

10

Date

TSP µg/m³

Kowhai, Kingsland

Gavin Street, Penrose

Kelly Street, Mt Eden

MoH Guideline = 60µg/m³


Figure 4:
Christchurch TSP 7-day average January–December 2010
[image: image12.emf]0

10

20

30

40

50

60

70

80

90

100

Jan 10 Mar 10 May 10 Jul 10 Sep 10 Nov 10

Date

TSP µg/m³

MOH Guideline = 60 µg/m³


Figure 5:
Christchurch TSP 7-day average 2002–2010
[image: image13.emf]0

10

20

30

40

50

60

70

80

90

100

Dec-02 Jun-03 Dec-03 Jun-04 Dec-04 Jun-05 Dec-05 Jun-06 Dec-06 Jun-07 Dec-07 Jun-08 Dec-08 Jun-09 Dec-09 Jun-10 Dec-10

Date

TSP µg/m³

MoH Guideline = 60µg/m³


7.8
Lead (Pb) June–August 2010
Lead is measured from seven-day averaged TSP samples to derive a 3-month average at Gavin Street, Penrose and Coles Place, St Albans. The results are described in the table below. Figure 6 provides moving 3-month averaged lead data between January 1996 and September 2000 when lead monitoring was performed on a monthly basis. From this point lead was monitored over a 3-month period (June to August) annually.

	Site
	June 2010
average (µg/m3)
	July 2010
average (µg/m3)
	August 2010
average (µg/m3)
	Winter 2010
average (µg/m3)

	Gavin Street, Penrose
	0.008
	0.008
	0.005
	0.007

	Coles Place, St Albans
	0.023
	0.029
	0.011
	0.021


No site exceeded the 3-month average guideline for lead (0.2 µg/m3).
Figure 6:
MfE lead 3-month average results 1996–2010
[image: image14.emf]0

0.05

0.1

0.15

0.2

1996 1997 1998 1999 2000 2001 2002 2003 2004 2005 2006 2007 2008 2009 2010

Date

Lead µg/m³

St Albans

Penrose

Mt Eden

Kowhai

AAQG 2002 = 0.2 µg/m

3


Figure 7:
MfE Burnside CO 1-hour fixed average January–December 2010

[image: image15.png]a

E

gy

10

0
Jan1o

Feb10

Mar10

Apri0

May10

Junto

10
Date

Augio

Sept0

oct10

Novio

Dec1o





Figure 8:
MfE Burnside CO 1-hour fixed average 2003–2010
[image: image16.png]a

E

gy

10

Jano3

Jan0s

Janos

Janos

Jan07
Date

Jan0s

Janos

Jan1o





Figure 9:
MfE Burnside CO 8-hour rolling average January–December 2010

[image: image17.png]2

10

majm®

o

Jani0 Febid Marid Aprid Mayio Jumid  Julio

2uglo Septo 0o Movio
Date

Decto





Figure 10:
MfE Burnside CO 8-hour rolling average 2003–2010
[image: image18.png]2

10

majm®

2an03 Jan0s Jam05 Jan06, Jan07 Janoe, Jan05 Jan0
Date




Figure 11:
MfE Burnside CO annual average 2003–2010
[image: image19.emf]0.0

0.1

0.2

0.3

0.4

0.5

0.6

0.7

0.8

0.9

1.0

2003 2004 2005 2006 2007 2008 2009 2010

Date

mg/m³


Figure 12:
MfE Penrose NO2 1-hour fixed average January–December 2010
[image: image20.png]240
220
200
180
160
140

E120

ST RTY T

Mario  Apri0  Maylo Juni0  Juil0  Augld  Sepl0 OCti0  Novid Decid
Date

Janio  Febi0





Figure 13:
MfE Penrose NO2 1-hour fixed average 1997–2010
[image: image21.png]200
220
200
180
160

L0

120

100

50

&0

w0

20

o

Jand7 JanSs Jan9e Jan00 Jan0L Jan02 Jan03 Jan04 Jan0S Jan0 Jan07 Jan0S Jan0S Jani0
Date




Figure 14:
MfE Penrose NO2 24-hour fixed average January–December 2010
[image: image22.png]110

100

10
0

Jani0  Febl0 Mario  Apri0  Mayio Juni0  Jul0  Auglo  Sepl0 Octid Novi0  Dect0
Date




Figure 15:
MfE Penrose NO2 24-hour fixed average 1 January 1997–2010
[image: image23.png]110
100

ANAMAMAMALA

JanS7 Janss JanSs Jan00 Jan0L Jan02 Jan03 Jan0é Jan0S Jan0 Jan07 Jan08 Jan0S Janid
Date





Figure 16:
MfE Penrose NO2 annual average 1997–2010
[image: image24.emf]0

5

10

15

20

25

30

1997 1998 1999 2000 2001 2002 2003 2004 2005 2006 2007 2008 2009 2010

Date

µg/m³


Figure 17:
MfE Penrose NO2 and NO 1-hour fixed average January–December 2010
[image: image25.png]bafms

800

700

500

s00

300

200

100

0
Jan1o

Feb10

Mar1o

Aprio

May10

Junto

10
Date

Augio

Sep10

oct10

Nov1o

Dec1o




Figure 18:
MfE Penrose NO2 and NO 24-hour fixed average January–December 2010
[image: image26.png]200
220
200
180
160

L0

120

100
50
&0
w0
20

o

Jani0  Febl0 Mario  Apri0  Maylo Juni0  Jull0  Augl0  Sepl0  Octid Novi0  Dect0
Date




Figure 19:
MfE Penrose NO 1-hour fixed average 1997–2010
[image: image27.png]bafms

1200

1000

800

400

0
Jans7

Janss

Janss

Janoo

Jano1

Jan02

Jano3

Jan0s
Date

Janos

Janos

Jan07

Jan0s

Janos

Jan1o




Figure 20:
MfE Penrose NO 24-hour fixed average 1997–2010
[image: image28.png]300

s0

0
Jand7 JanSs JanSs Jan00 Jan0L Jan02 Jan03 Jan04 Jan0S Jan0 Jan07 Jan08 Jan0S Janid
Date





Figure 21:
MfE Penrose NO annual average 1997–2010
[image: image29.emf]0

10

20

30

40

50

60

1997 1998 1999 2000 2001 2002 2003 2004 2005 2006 2007 2008 2009 2010

Date

µg/m³


Figure 22:
MfE Burnside NO2 1-hour fixed average January–December 2010
[image: image30.png]220

200

180
160
140

20

S100
50
&0
w0
20

0
Jani0  Febl0 Marid  Apri0  Maylo Juni0  Jull0  Augld  Sepl0 OCti0  Novid Decid
Date





Figure 23:
MfE Burnside NO2 1-hour fixed average 2003–2010
[image: image31.png]220
200
180
160
140

21
S100

50

50





Figure 24:
MfE Burnside NO2 24-hour fixed average January–December 2010
[image: image32.png]110
100

10
0

et VA o,

Jani0  Febl0 Mario  Apri0  Maylo Juni0  Jull0  Augl0  Sepl0  Octi0 Novi0  Dect0

Date




Figure 25:
MfE Burnside NO2 24-hour fixed average 2003–2010
[image: image33.png]110
100

Janos Janos Janos Jan07 Jan0s Janos Jan1o
Date




Figure 26:
MfE Burnside NO2 annual average 2003–2010
[image: image34.wmf]0

5

10

15

20

25

30

2003

2004

2005

2006

2007

2008

2009

2010

Date

µg/m³

 


Figure 27:
MfE Burnside NO2 and NO 1-hour fixed average January–December 2010

[image: image35.png]bafms

800

700

500

s00

300

200

100

0
Jan1o

Feb10

Mar10

Aprio

May10

Junto

10
Date

Augio

Sept0

oct10

Nov1o

Dec1o




Figure 28:
MfE Burnside NO2 and NO 24-hour fixed average January–December 2010
[image: image36.png]10 Auglo  Sept0  Octio  Novid  Decid
Date

Jri0  Febi0 Marlo  Aprio  Mayio  Jumn




Figure 29:
MfE Burnside NO 1-hour fixed average 2003–2010

[image: image37.png]1200

1100
1000
s00
800
L, 700
£ o0
= s00
400
300
200
100

Jano3 Jan0s Janos Janos Jan07 Jan0s Janos Jan1o
Date




Figure 30:
MfE Burnside NO 24-hour fixed average 2003–2010
[image: image38.png]bgfms

300

250

200

100

s0

Jano3

Jan0s

Janos

Janos

Jan07
Date

Jan0s

Janos

Jan1o




Figure 31:
MfE Burnside NO annual average 2003–2010
[image: image39.emf]0

10

20

30

40

50

60

2003 2004 2005 2006 2007 2008 2009 2010

Date

µg/m³


Figure 32:
MfE Penrose SO2 1-hour fixed average January–December 2010
[image: image40.png]400

300

vafm

100

0
Jan1o

Feb10

Mar1o

Apri0

May10

Junto

10
Date

Augio

Sep10

oct10

Novio

Dec1o





Figure 33:
MfE Penrose SO2 1-hour fixed average 2003–2010

[image: image41.png]400

300

vafm

100

ol
Jano3

Jan0s

Janos

Janos

Jan07
Date

Jan0s

Janos

Jan1o





Figure 34:
MfE Penrose SO2 24-hour fixed average January–December 2010
[image: image42.png]bafms

140

120

100

2

20

o4
Jan1o

Feb10

Mar1o

Aprio

May10

Junto

10
Date

Augio

Sep10

oct10

Nov1o

Decto





Figure 35:
MfE Penrose SO2 24-hour fixed average 2003–2010

[image: image43.png]140

120

100

pg/m=
2 @
g 8





Figure 36:
MfE Penrose SO2 annual average 2003–2010

[image: image44.emf]0

1

2

3

4

5

6

7

8

9

10

2003 2004 2005 2006 2007 2008 2009 2010

Date

µg/m³


Figure 37:
MfE Burnside SO2 1-hour fixed average January–December 2010
[image: image45.png]400

300

vafm

100

o S Lot o B b et b s

Mayl0 Junlo  Jull0  Auglo  Sepld
Date

Octlo  Novio Decto




Figure 38:
MfE Burnside SO2 1-hour fixed average 2003–2010

[image: image46.png]400

300

vafm

100

Jano3

Jan0s

Janos

Janos

Jan07
Date

Jan0s

Janos

Jan1o





Figure 39:
MfE Burnside SO2 24-hour fixed average January–December 2010
[image: image47.png]bafms

140

120

100

@

20

0

A

s itanorbnon o M Ao

A s

Jan1o

Feb10

Mario  Aprio  Mayio Junto  Jul1o  Augio
Date

Sep10

oct10

Novio

Decto





Figure 40:
MfE Burnside SO2 24-hour fixed average 2003–2010

[image: image48.png]bafms

140

120

100

2





Figure 41:
MfE Burnside SO2 annual average 2003–2010
[image: image49.emf]0

1

2

3

4

5

6

7

8

9

10

2003 2004 2005 2006 2007 2008 2009 2010

Date

µg/m³


Figure 42:
MfE Penrose PM10 24-hour fixed average January–December 2010

[image: image50.png]10
0

T

Jani0  Febl0 Mario  Apri0  Mayio Juni0  Jul0  Auglo  Sepl0 Octid Novi0  Dect0

Date




Figure 43:
MfE Penrose PM10 24-hour fixed average 2003–2010
[image: image51.png]140

120

100

20

Jano3

Jan0s

Janos

Jans,

Jan07
Date

Jan0s

Janos

Jan1o





Figure 44:
MfE Penrose PM10 annual average 2003–2010

[image: image52.emf]0

5

10

15

20

25

30

2003 2004 2005 2006 2007 2008 2009 2010

Date

µg/m³


Figure 45:
MfE Burnside PM10 24-hour fixed average January–December 2010
[image: image53.png]10

Jai

*NMMWWNJW\’WWW\WWWWM

N0 Febl0 Marlo  Apri0  Mayid Junid  Juil0  Auglo  Sepl0 Octio  Novid  Decid
Date





Figure 46:
MfE Burnside PM10 24-hour fixed average 2003–2010
[image: image54.png]okl

Jano3

Janos Janos Janos, Jan07 Jan0s Janos Jan1o
Date




Figure 47:
MfE Burnside PM10 annual average 2003–2010

[image: image55.emf]0

10

20

30

40

50

60

2003 2004 2005 2006 2007 2008 2009 2010

Date

µg/m³


7.9
Analysis of exceedences

The only exceedences recorded during 2010 were at Greers Road, Burnside, for PM10 daily averages. All the other parameters monitored at all sites were below the National Environmental Standards (NES) for Air Quality.

7.9.1
Exceedences at Greers Road, Burnside

Time of year

The PM10 standard of 50 µg/m3 was exceeded on eight days in 2010 during the cooler months from June to July. During July there were three consecutive days that incurred exceedences
(27–29 July).

Time of day

The typical diurnal trend in PM10 during winter is shown in Figure 48. Here PM10 hourly averages for those days where the daily PM10 average exceeded 50 µg/m3 are plotted against time of day. From Figure 48, the biggest contributions to PM10 levels were between the hours of 19:00 and 04:00. This would suggest contributions from wood burning for home heating. On 21 June, PM10 concentrations remained higher at night compared to the other exceedence days. On 27 to 29 July elevated PM10 levels occur from the early morning (01:00).
Temperature

Figures 49, 50 and 51 plots PM10 hourly averages for those days where the daily PM10 average exceeded 50 µg/m3 against hourly average temperature measured at 1.5 m, and 10 m, and the difference between the two temperature heights. From Figure 49, the biggest contributions to PM10 levels were when temperature measured at 1.5 m was below 6 °C. From Figure 50, the biggest contributions to PM10 levels were when temperature measured at 10 m was below 7 °C. On 11 July, both temperatures were lower compared to the other exceedence days. From Figure 51, the biggest contributions to PM10 were when temperature difference between 10 and 1.5 m was greater than zero, ie, when temperature inversion conditions prevailed. This is consistent with the trapping of pollutants and subsequent higher concentrations expected during temperature inversion conditions.
Relative humidity

Figure 52 plots PM10 hourly averages for those days where the daily PM10 average exceeded 50 µg/m3 against hourly average relative humidity (RH). From Figure 52, PM10 levels began to rise when RH was above 75 per cent and peaked at around 90 per cent. This is consistent with diurnal RH patterns throughout the day.

Wind direction

Figure 53 plots PM10 hourly averages for those days where the daily PM10 average exceeded 50 µg/m3 against hourly average wind directions. From Figure 53, no wind direction seems to contribute significantly higher concentrations than others. This is consistent with an area with wide diffuse sources of pollution as opposed to point/line source of pollution. There were very few data points in the sector 120 to 180 degrees as wind frequency from this sector is usually low.

Wind speed

Figure 54 plots PM10 hourly averages for those days where the daily PM10 average exceeded 50 µg/m3 against hourly average wind speeds. From Figure 54, contributions to PM10 began when wind speed was below 1.0 m/s and peaked at very low wind speed conditions. This is consistent with reduced dispersion under low wind speed conditions. 21 June and 27 July had higher wind speed compared to the other exceedence days. These winds increased from mid morning increasing dispersion and decreasing the PM10 concentrations.
NO
Figure 55 plots PM10 hourly averages for those days where the daily PM10 average exceeded 50 µg/m3 against hourly average NO. From Figure 55, a positive linear relationship seems to exist between PM10 and NO concentrations. This relationship is usually more defined at PM10 concentrations below 60 µg/m3 and NO concentrations below 80 µg/m3. This is expected as PM10 and NO is co-generated during the burning of wood and fossil fuels.

CO

Figure 56 plots PM10 hourly averages for those days where the daily PM10 average exceeded 50 µg/m3 against hourly average CO. From Figure 56, a positive linear relationship seems to exist between PM10 and CO concentrations. This relationship is usually more defined at PM10 concentrations below 60 µg/m3 and CO concentrations below 2.5 mg/m3. This is expected as PM10 and CO is co-generated during the burning of wood and fossil fuels.
SO2
Figure 57 plots PM10 hourly averages for those days where the daily PM10 average exceeded 50 µg/m3 against hourly average SO2. From Figure 57, the relationship between PM10 and SO2 concentrations seems to depend on the exceedence day. On a few days there seems to be a positive trend although this is probably due to meteorological conditions.
Conclusions

From the comparisons of available meteorological parameters and other pollutants as discussed above, it can be concluded that PM10 exceedences were likely due to home heating during the colder months. This is observed especially in the evenings and early mornings, when PM10 levels are worsened by temperature inversions trapping the pollutants and low wind speed conditions preventing effective dispersion of pollutants.

Figure 48:
MfE Burnside, PM10 vs time of day

[image: image56.emf]0

50

100

150

200

250

300

0 2 4 6 8 10 12 14 16 18 20 22 24

Time of day

PM

10

 µg/m³

21/06/2010

1/07/2010

3/07/2010

11/07/2010

13/07/2010

27/07/2010

28/07/2010

29/07/2010

All days

average


Figure 49:
MfE Burnside, PM10 vs 1.5 m temperature

[image: image57.emf]0

50

100

150

200

250

300

-10 -8 -6 -4 -2 0 2 4 6 8 10 12 14 16 18 20

Temperature (°C) at 1.5m

PM

10

 µg/m³

21/06/2010

1/07/2010

3/07/2010

11/07/2010

13/07/2010

27/07/2010

28/07/2010

29/07/2010


Figure 50:
MfE Burnside, PM10 vs 10 m temperature

[image: image58.emf]0

50

100

150

200

250

300

-10 -8 -6 -4 -2 0 2 4 6 8 10 12 14 16 18 20

Temperature (°C)  at 10m

PM

10

 µg/m³

21/06/2010

1/07/2010

3/07/2010

11/07/2010

13/07/2010

27/07/2010

28/07/2010

29/07/2010


Figure 51:
MfE Burnside, PM10 vs temperature diff 10–1.5m

[image: image59.emf]0

50

100

150

200

250

300

-4 -3 -2 -1 0 1 2 3 4

Temperature difference (°C)

PM

10

 µg/m³

21/06/2010

1/07/2010

3/07/2010

11/07/2010

13/07/2010

27/07/2010

28/07/2010

29/07/2010


Figure 52:
MfE Burnside, PM10 vs. relative humidity
[image: image60.emf]0

50

100

150

200

250

300

0 10 20 30 40 50 60 70 80 90 100

RH (%)

PM

10

 µg/m³

21/06/2010

1/07/2010

3/07/2010

11/07/2010

13/07/2010

27/07/2010

28/07/2010

29/07/2010


Figure 53:
MfE Burnside, PM10 vs wind direction

[image: image61.emf]0

50

100

150

200

250

300

0 30 60 90 120 150 180 210 240 270 300 330 360

Wind direction

PM

10

 µg/m³

21/06/2010

1/07/2010

3/07/2010

11/07/2010

13/07/2010

27/07/2010

28/07/2010

29/07/2010


Figure 54:
MfE Burnside, PM10 vs wind speed
[image: image62.emf]0

50

100

150

200

250

300

0 0.5 1 1.5 2 2.5 3 3.5 4 4.5 5 5.5 6 6.5 7

Wind speed (m/s)

PM

10

 µg/m³

21/06/2010

1/07/2010

3/07/2010

11/07/2010

13/07/2010

27/07/2010

28/07/2010

29/07/2010


Figure 55:
MfE Burnside, PM10 vs NO
[image: image63.emf]0

50

100

150

200

250

300

0 50 100 150 200 250 300 350 400 450 500

NO (µg/m³)

PM

10

 µg/m³

21/06/2010

1/07/2010

3/07/2010

11/07/2010

13/07/2010

27/07/2010

28/07/2010

29/07/2010


Figure 56:
MfE Burnside, PM10 vs CO
[image: image64.emf]0

50

100

150

200

250

300

0 1 2 3 4 5 6 7 8 9

CO (mg/m³)

PM

10

 µg/m³

21/06/2010

1/07/2010

3/07/2010

11/07/2010

13/07/2010

27/07/2010

28/07/2010

29/07/2010


Figure 57:
MfE Burnside, PM10 vs SO2
[image: image65.emf]0

50

100

150

200

250

300

0 5 10 15 20 25

SO

2

 (µg/m³)

PM

10

 µg/m³

21/06/2010

1/07/2010

3/07/2010

11/07/2010

13/07/2010

27/07/2010

28/07/2010

29/07/2010





































































































































































































MfE Gavin Street, Penrose





AAQG = 20 µg/m³





NES = 50 µg/m³





NES = 50 µg/m³





AAQG = 20 µg/m³





NES = 50 µg/m³





NES = 50 µg/m³





AAQG = 120 µg/m³





AAQG = 120 µg/m³





NES = 350 µg/m³





NES = 350 µg/m³





AAQG = 120 µg/m³





AAQG = 120 µg/m³





NES = 350 µg/m³





NES = 350 µg/m³





AAQG = 100 µg/m³





AAQG = 100 µg/m³





NES = 200 µg/m³





AAQG = 100 µg/m³





AAQG = 100 µg/m³





NES = 200 µg/m³





NES = 200 µg/m³





NES = 10 mg/m³





NES = 10 mg/m³





AAQG = 30 mg/m³








AAQG = 30 mg/m³














NES = 200 µg/m³








�	Ministry for the Environment. 2005. Updated Users Guide to Resource Management (National Environmental Standards Relating to Certain Air Pollutants, Dioxins and Other Toxics) Regulations 2004. Wellington: Ministry for the Environment.


�	Schwela DH. 1999. Public Health and the air management information system (AMIS). Epidemiology, 10(5), 647-655.


�	Ministry for the Environment. 2007. Environment New Zealand 2007. Wellington: Ministry for the Environment.


�	Ministry for the Environment. 2003. Emission Inventories for CO, NOx, SO2, ozone, benzene and benzo(a)pyrene in New Zealand. http://www.mfe.govt.nz/publications/air/air-quality-tech-report-44-nov03/index.html


�	Ministry for the Environment. 2002. Ambient Air Quality guidelines – 2002 Update. Wellington: Ministry for the Environment.


�	Source: http://www.mfe.govt.nz/laws/standards/air-quality-standards.html.





[image: image67.jpg]

